+32 16 90 39 04

https://xeryon.com

Diestsesteenweg 692, 3010 Leuven, Belgium

XD-OEM Controller manual – 3.2

Image for illustration only. May not be an exact representation of the product.

The XD-OEM is a versatile industrial piezo controller, supporting a wide variety of use cases and  communication options for our actuators and stages. Designed for OEM integration, both single-axis  and multi-axis (through EtherCAT) is supported. A Windows GUI is included for plug-and play testing.

XD-OEM Controller manual – 3.2 Last updated: 28/08/2025. All specifications are subject to change without prior notice.

CONTENT

1. Key Specifications ........................................................................................................................................... 3 2. Getting started ................................................................................................................................................ 4 3. Control of an actuator .................................................................................................................................... 6 4. Communication using USB / UART ............................................................................................................... 7

4.1. Format for instructions........................................................................................................................ 7 4.2. List of commands ................................................................................................................................. 8 4.3. Feedback from controller..................................................................................................................13

5. Communication using GPIO ........................................................................................................................15 5.1. Pulse and direction mode .................................................................................................................15 5.2. Encoder-like signal (A quad B input)................................................................................................15 5.3. PWM control........................................................................................................................................16 5.4. Analog controls...................................................................................................................................16

6. Communication using EtherCAT .................................................................................................................17 7. Connections / pin layouts ............................................................................................................................18 7.1. Connection to actuator......................................................................................................................18 7.2. Connection to stage ...........................................................................................................................18 7.3. Connection to GPIO ...........................................................................................................................19 8. Tuning.............................................................................................................................................................20 9. Windows GUI .................................................................................................................................................22 9.1. Required files ......................................................................................................................................22 9.2. Commands for the Windows GUI demo program.........................................................................22 10. Python, MATLAB, C++ and LabVIEW ...........................................................................................................25 11. Support...........................................................................................................................................................26

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

1. KEY SPECIFICATIONS

Compatible products XLA series, XLS series, XRT series Number of axes Single-axis (multi-axis through EtherCAT, see Ch. 6) Power supply 12 – 48 VDC (screw terminal), 10 W per controller Power supply (XLA-10 only) 24 – 48 VDC (screw terminal), 20 W per controller Temperature range -30°C to +70°C

Dimensions 111 x 25 x 13.2 mm

Connector to actuator / stage ZIF | Molex

Control types Closed loop, open loop, hybrid Key communication options USB, UART, EtherCAT, GPIO, … Software support Windows GUI, C++, Python, TwinCAT, LABView Motion profile Point-to-point trapezoidal

Encoder compatibility Optical, incremental

Supported resolutions Up to 78 nm (linear), up to 3 µrad (rotary) Error detection Overcurrent, thermal, short-circuit, …

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

2. GETTING STARTED

All controllers are delivered with a power adapter, USB cable and USB stick with software to control  actuators and stages (from now on: ‘actuator(s)’) for easy plug-and-play use. To get started, plug the  actuator connector into the XD-OEM. Connect the XD-OEM to a computer with the provided USB cable  and open up the Windows GUI, press SCAN or MOVE and the actuator will start moving.

A quick overview of the Windows GUI:

To control an actuator:

1. Connect the correct COM port.

2. Load settings_default.txt from USB stick.

3. Press “Find Index”.

The actuator will start moving and look for the encoder index.

4. You can start controlling the actuator:

a. SCAN: move with continuous speed.

b. STEP: move with specified step size.

c. 0: move to home position.

Using the XD-OEM controller, a Xeryon actuator can be used to create a wide variety of motion  patterns, such as:

• Short- or long-range motion, high- or low-speed motion

• Repetitive or dynamic motion, smooth continuous motion

• Set-and-forget applications, controlled damping (soft landing) motion

• A combination of the above, and more

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

Tuning

Xeryon actuators work based on ultrasonic piezo technology. This technology is frequency-based, and  differs slightly from both a) electromagnetic motors and b) other piezo systems. As a result, motion  characteristics may differ from conventional motion systems. Several parameters can be used for  finetuning motion behaviour. Motion trajectory is split into two zones: bulk move (or ‘scanning’) and  final positioning. When experiencing unexpected behaviour, the following simple tuning procedures  can be followed (Settings -> Edit). In-depth tuning info can be found in chapter 9.

Behaviour Final positioning (zone 1) Bulk move / scanning (zone 2)

Vibration Lower proportional factor Lower proportional factor

Adjust frequency

Lower deceleration

-

Noise Adjust frequency Overshoot Lower deceleration Force too low Lower frequency Inaccurate landing Lower positioning tolerance -

Note: for vertical or inclined use with higher payloads, advanced tuning may be necessary. For assistance,  please contact support@xeryon.com. More information about the Windows GUI: ch. 9.

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

3. CONTROL OF AN ACTUATOR

The controllers can be controlled using various methods:

• Using USB / UART see section 4

• Using GPIO see section 5

• Using EtherCAT see section 6

For nanosecond multi-axis use, EtherCAT is recommended. For fast microsecond pulses: GPIO

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

4. COMMUNICATION USING USB / UART

A host computer or controller can communicate with the XD-OEM via the USB configured as a virtual  COM port. The baud rate is automatically detected by the controller and can be up to 115200. The  protocol uses 8 data bits, 1 stop bit, no parity bit, no handshaking. This section consists of three parts:

• 4.1 Format for instructions

• 4.2 List of commands

• 4.3 Feedback from controller

Note: To enable UART, send UART=9600 command and use the following pins: Ground (pin 2), 5V reference (5), Tx (12), Rx (18).

4.1. Format for instructions

A command line consists of maximum 16 characters followed by a ‘new line’ character (ASCII code 10).  The command has the following fixed format: X:DPOS=-12345678

• 1 character defining the axis, followed by a colon.

• 4 characters for the command.

• ‘=’ sign separating the command from the corresponding value.

• Optional sign.

• Decimal value up to 8 decimal places (9 if the sign is omitted).

• Maximum total of 16 characters.

The characters have to be sent from left to right, in the example above starting with ‘X’ and ending with  ‘8’. The command tags are in upper case. The instruction should be terminated with a ‘new line’  character (ASCII code 10). The driver processes the instruction immediately after receiving this ‘new line’  character.

Some instructions such as ‘ZERO’ and ‘RSET’ require no value. In that case, it is sufficient to send only  the command itself, e.g. ‘ZERO’ followed by the ‘new line’ character.

Value range

There are 9 characters reserved for the value including its sign. For signed values 8 decimal places are  available, giving a range from -99 999 999 to +99 999 999. For positive numbers, the ‘+’ sign can be  omitted, increasing the positive range to 999 999 999. No spaces, commas or periods should be added  to the numbers. Only integers are allowed.

• X:DPOS=-99999999

• X:DPOS=+99999999

• X:DPOS=999999999

Request a value

To request the value of a certain setting, put ‘=?’ after the parameter for which you want to know the  value, e.g. EPOS=? gives the controller a request for the encoder position. FREQ=? asks the controller  for the current excitation frequency. This works best with INFO=0, otherwise the reply disappears in the  constant flood of feedback data.

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

Units are as follows:

Type XLA / XLS series XRT series Resolution

Time, delays ms 1 ms

| encoder units |  |
| --- | --- |
| µm/s | deg/s (*) |

Target position, step size 1 encoder increment Speed

Frequency Hz 1 Hz

(*) Conversion factor of 100 required: e.g. enter SSPD=10000 for 100 deg/s.

4.2. List of commands

This is a list of all possible commands that can be sent and their use. This part is divided into: I. Motion

II. Settings handling

III. Communication

IV. Tuning

V. Signal (advanced)

VI. Directional settings (advanced)

VII. Trigger outputs (advanced)

I. MOTION

Command Range Explanation

0, 1

1 bit

-

26 bits

26 bits

-1,0,1

INDX Find the index. A value of 0 or 1 indicates the initial search direction. The  controller sets off in the specified direction to search. When a mechanical

limit is reached (detected by position error > ILIM) it reverses the search

direction. After finding index, actuator is positioned at the index position.

INDA Automatic detection of encoder index.

HOME Go to the home position. This equals DPOS=0.

DPOS Set target position. Closed-loop control is used to reach and maintain the  new position. The position is expressed in encoder units. Positive and

negative values are allowed within the range of the actuator.

STEP Move relative to the current position, over a specified distance. When  already in closed loop, the current desired position is used as a reference.

When before in open loop, the actual position (encoder value) is used as a

reference. The command value specifies the step size in encoder

increments. Positive values send the actuator towards higher encoder

values, negative values send the actuator towards lower encoder values.

Closed-loop control is used to reach and maintain the new position.

SCAN Continuously move with fixed speed. The speed is maintained by closed loop control. A positive number sends the actuator towards increasing

encoder values, a negative number sends the actuator towards decreasing

encoder values. A zero value stops the actuator.

SSPD 24 bits Set speed. Used as scanning speed (SCAN command) and as target speed  towards the next target position (DPOS and STEP). Unit is 1 µm/s or 0.01

deg/s. Default: 10000 (10 mm/s or 100 deg/s).

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

I. MOTION

Command Range Explanation

24 bits

16 bits

16 bits

26 bits

26 bits

-1,0,1

16 bits

16 bits

-

-

1 bit

1 bit

ISPD Set the speed which is used while searching the index. Unit is 1 µm/s or 0.01 deg/s.

ACCE Set acceleration for speed profile. Expressed in m/s2. Default value: 65500.

DECE Set deceleration for speed profile, when approaching target position.  Default / maximum value: 65500 for XD-OEM controller.

LLIM Set low-side soft end stop. Expressed in encoder units. HLIM Set high-side soft end stop. Expressed in encoder units.

MOVE Continuously move in open loop. Phase and amplitude influence the speed,  but speed is not controlled. A positive number sends the actuator towards

increasing encoder values, a negative number sends the actuator towards

decreasing encoder values. A zero value stops the actuator.

PHAS Set the phase offset between the excitation signals. Can be used to control  the speed in open loop. Input values 0-65535 correspond to a phase shift of  0-360°. Around 16384 the phase corresponds with a MOVE=1 direction,

around 49152 (=-16384) it corresponds to a MOVE=-1 direction.

AMPL Set amplitude for open-loop piezo excitation signals. STOP Stop the actuator.

CONT Continue movement after a stop command.

ENBL Enable drive. Or: If an actuator is in error mode, you can recover from that  by sending “ENBL=1”. This enables the actuator again.

ENBR Enable drive on power-up automatically.

FILE Filters fast variations from the encoder values (noise).  Default value: 0, sets filter to a minimum.

FILG Extracts the spikes from the raw encoder inputs. Defines the minimum  pulse width that may be responded to. Default value: 0

INTF Controller integration factor. Default value: 1. Minimum = 1

DTIM Determines how long (in ms) the motor must be turned off before it can be  turned on again. Ensures that the actuator is not over-excited when landing

by switching it on and off too quickly. Default value: 0

II. MANAGE SETTINGS

Command Range Explanation

-

-

-

0-1

RSET Reset the driver. All piezo signals go to zero and settings are reset to their  saved value.

LOAD Load settings from memory.

SAVE Save settings to memory.

BLCK Blocked mode. When blocking mode is enabled (1), the ENBL command  should be send to the controller if an error occurs. If blocking mode is

disabled (0), the controllers executes the next command even when an

error occurs.

TEST 0-1 Test LED indicators. TEST=1 switches all indicators on. TEST=0 brings them  back to their function.

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

III. COMMUNICATION

Command Range Explanation

4 bits

0-76800

0-13

16 bits

1-65535

INFO Select type of info to be transmitted from the controller to the master (PC). 0: Stop broadcasting info automatically.

1: SRNO, SOFT, MODEL, STAT, SYNC

2: SRNO, SOFT, MODEL, STAT, FREQ, SYNC, EPOS, DPOS, REQUESTED

PARAMETER*, TIME

3: EPOS, DPOS, STAT

4: EPOS, STAT, DPOS, TIME

5: STAT, FREQ, EPOS, DPOS, REQUESTED PARAMETER*, TIME

6: REQUESTED PARAMETER*

7: EPOS, STAT

Default: 2

e.g. INFO=7 will alternatingly send EPOS & STAT values.

* REQUESTED VALUE, when a parameter value is requested (e.g. by sending

FREQ=?) the value for the parameter will be returned here.

UART Set UART baud rate. To switch UART off: UART=0. To set the baud rate to  9600: UART=9600. The maximum baud rate is 76800. When UART is off, the

UART can only be restarted by sending the UART command via USB.

GPIO Select the preferred input mode using the GPIO command. For more details see section 5.

| GPIO=0 | Control via IO pins switched off. The controller will only react to text commands sent through USB or UART. |
| --- | --- |
| GPIO=2 | Pulse and direction mode, with direction & enable pins. |
| GPIO=3 | Pulse and direction mode, with forward & backward pins. |
| GPIO=4 | A quad B input mode. |
| GPIO=8 | PWM control, with direction & enable pins. |
| GPIO=9 | PWM control, with forward & backward pins. |
| GPIO=12 | Analog control, with direction & enable pins. |
| GPIO=13 | Analog control, with forward & backward pins. |

STPS Select the step size for each STEP pulse. The step size is expressed in  encoder units. Default: 1 encoder unit.

POLI Set polling interval. Specifies the time between data updates. Also defines  the polling interval of the analog IO. The interval is expressed in

milliseconds. The default value is 97 (97 ms).

DLAY 16 bit Sets the delay between the moment the actuator reaches its target position  and the moment the ‘position reached’ flag is raised. Expressed in

milliseconds. Default: 1 (ms).

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

IV. TUNING

Command Range Explanation

24 bits

24 bits

16 bits

16 bits

26 bits

26 bits

1 bit

20 bits

26 bits

26 bits

16 bits

16 bits

16 bits

16 bits

FREQ Set the frequency of the excitation signals for zone 1. Unit is Hz.

FRQ2 Set the frequency of the excitation signals for zone 2. Unit is Hz. Also used  for scanning.

PROP Proportional control factor for zone 1.

PRO2 Proportional control factor for zone 2.

ZON1 Width of zone 1: Expressed in mm.

ZON2 Width of zone 2: Expressed in mm.

DUCO Amplitude is used in closed loop if set to 1. If set to 0, a fixed amplitude of  50% is used. Default: 1.

ELIM ELIM (error limit) sets the maximum following error. When the following  error exceeds the value set bij ELIM, then the controller goes in safe mode

and the motor signals are switched off. Recovery: RSET or ENBL=1.

This error may be triggered when trying to move beyond the physical limits

of the actuators, or by setting too high a speed. Do not forget to first find

the index position (INDX command) to avoid that the actuators travels

beyond the end stops and triggers this error. Default: 10000.

ILIM Sets the following error at which the index search algorithm reverses  direction. This influences the time the actuator stalls at the end position

during an index search.

SLIM Error saturation limit

PTOL Position tolerance. When the actuator is within +/- position tolerance of the  desired position, the control is switched off and the ‘position reached’ flag is  raised.

Values are expressed in encoder units and should be in the range 0 –

65535. The range is applied symmetrically with respect to positive and

negative position errors. e.g. PTOL=2 allows s errors between -2 and +2

encoder units. Default: 2. See also TOUT and PTO2.

PTO2 Second position tolerance, similar to PTOL. Comes into action if first  position tolerance PTOL fails within a timeout time TOUT. The default value

is 10.

TOUT Set timeout time. To avoid that the actuator keeps vibrating indefinitely  around the desired position without ‘landing’, a timeout time can be set.

The timer starts when the actuator is near the desired position, within a

distance of +/- PTO2. After passing the timeout time, PTO2 becomes the

new position tolerance.

The time is expressed in milliseconds. The default value is 1000 (ms).

TOU2 This defines a safety timeout. When the motor is on for a time longer that  the value set by TOU2, then the controller goes in safe mode and the motor

signals are switched off. Recovery: RSET or ENBL=1 depending on BLCK.

Status bit #18 goes up when this timeout is triggered.

TOU2=0 disables this timeout. Any other value sets the timeout time in

seconds. Minimum value: 1 second. Maximum value: 65535 seconds.

TOU3 16 bits This defines another safety timeout. When the actuator is trying to “land” to  a specific position for longer than the value set by TOU3, then the controller

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

IV. TUNING

Command Range Explanation

goes in safe mode. The motor signals are switched off. Expressed in  milliseconds.

ENCR 1 bit Reset the encoder by sending “ENCR=1”.

V. SIGNAL SHAPING (ADVANCED USE)

Command Range Explanation

0-3

0-1

-

16 bits

16 bits

16 bit

12 bits

12 bits

8 bits

ENBL Enable amplifiers. Bit 0 is for piezo signal 1, bit 1 for piezo signal 2. ENBL=3  enables both amplifiers, ENBL=0 disables both amplifiers. ENBL=1 enables

only amplifier 1, ENBL=2 enables only amplifier 2.

ENBR Enable amplifiers automatically upon controller start-up. ZERO Force the piezo signals to zero volt.

MAMP Set maximum amplitude. The piezo excitation signals are limited to the  corresponding voltages. MAMP=65535 sets them to the maximum voltage

of 45 V. MAMP=36400 sets the maximum to 25 V. The relation is linear.

MIMP Set minimum amplitude for piezo excitation signals. See MAMP for values.

PHAC Phase correction. Corrects an imbalance in the motor. Such imbalance may  cause a rattling or scratching noise when the actuator moves at low speed.

Practical values are in the range of a few 1000, positive or negative.

Default: 0 (no correction)

OFSA Offset on the piezo signals on piezo phase 1. OFSA=4095 corresponds to full  scale (45 V), OFSA=0 produces no offset. The relation is linear.

OFSB Similar to OFSB, but for piezo phase 2.

FILP Filter speed for phase of piezo excitation signals. Default value: 1. Max.  value: 255.

FILA 8 bits Filter speed for amplitude of piezo excitation signals. Default value: 1. Max.  value: 255.

VI. DIRECTIONAL SETTINGS (ADVANCED USE)

Command Range Explanation

1 bit

32 bits

1 bit

ENCD Set the encoder direction. Set the counting direction with respect to the A/B  signals or sin/cos signals of the encoder. Flip this bit to swap left and right,

or clockwise and counter-clockwise. Default value is 0.

ENCO Sets the encoder offset: distance between the index position and the  desired zero position. In encoder units. Default value is 0.

ACTD Set the actuation direction. If not set correctly, the actuator will move away  from the desired position. Default value is 0.

PATH 1 bit For rotation stages only. Selects whether the stage will follow the shortest  path (PATH=1) to the target position or follow a linear approach, respecting

high to low or low to high (PATH=0). Default: 1 for rotation stages, 0 for

linear stages.

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

4.3. Feedback from controller

Status LEDs

Green LED Power on for controller board

Orange LED Motor on

Red LED Disabled (ENBL=1 to enable) OR error (position failure, amplifier fault, timeout)

Information is sent back from the Xeryon controller to the master in ASCII format. The format is as  follows:

1. Tag: Four characters describing the type of information

2. ‘=’ sign separating the command from the corresponding value

3. Signed value associated with that information (sign + 8 decimal places). The message is terminated  with a ‘new line’ character (ASCII code 10).

e.g. EPOS=12345678

Different types of information

The command INFO determines which information is sent back. See the example below (INFO=2). Tag Explanation

SRNO Serial number of the driver (hardware)

SOFT Software version installed on the driver. e.g. 20103 → 2.1.3

[TYPE] Type of motion device (XLA, XLS, XRT) and its resolution e.g. XLS1=312 STAT Status (see below)

FREQ Excitation frequency currently in use

SYNC Fixed value “12345678”. Can be used for debugging communication issues. EPOS Encoder position

DPOS Desired position

TIME Time stamp: resolution 0,1 ms

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

Meaning of STAT(US)

The Status contains 24 bits:

Status bit Name Explanation

Amplifiers enabled

End stop

Thermal protection 1

Thermal protection 2

Force zero

Motor on

Closed loop

Encoder at index

Encoder valid

Searching index

Position reached

Error compensation

Encoder error

Scanning

Left end stop

Right end stop

Error limit

Searching optimal frequency

Safety timeout triggered

EtherCAT acknowledge

Emergency stop

0 XRTA only: Amplifiers for phase 1 and 2 enabled

1 Actuator stopped by end stop

2 Amplifier for phase 1 or 3 in thermal protection. 3 Amplifier for phase 2 or 4 in thermal protection.

4 Motor signals are currently forced to zero.

5 The piezo motor is on.

6 The actuator is currently in closed loop control.

7 Indicates whether the actuator is positioned exactly at the encoder  index.

8 Indicates whether the encoder index has been passed and therefore  the encoder value reflects the absolute position, not the relative

position with respect to the startup position.

9 Indicates whether the actuator is currently searching the index  position.

10 Indicates whether the target position is reached (within tolerance  limits).

11 Error compensation is on.

12 Indicates an error produced by the encoder.

13 Indicates whether the actuator is in a scanning mode. 14 Indicates that the left end stop is passed.

15 Indicates that the right end stop is passed.

16 Indicates that the position error has reached the limit set by ELIM. This  can indicate a collision or mechanical limit (end of stroke).

17 The driver is searching for the optimal excitation frequency of the  piezo motor.

18 If this is set to “1”, then the safety timeout was triggered. See the  explanation of the command “TOU2” several pages back.

19 only used when control via EtherCAT

20 not used

21 Position fail If this is set to “1”, then the safety timeout TOU3 was triggered.

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

5. COMMUNICATION USING GPIO

There are 4 different configurations for the use of the digital and analog IO pins.  1. Control the position of the actuator using pulses. Each pulse does a step (size) in a certain direction. 2. Control the position of the actuator using an encoder-like signal.

3. Control the speed of the actuator using a PWM input signal.

4. Control the speed of the actuator using an analog input.

All of these methods are described more in detail below. To select a method, the command “GPIO” is  used. This can be found in section. This command can be sent over USB or UART. Corresponding pins  can be found in section 6.

The GPIO settings can be stored in memory using the “SAVE” command. That way, the GPIO mode will  become active at power-up without having to first send the GPIO command via USB, UART or any other  method.

The GPIO input and commands can be used together: the controller will react to both GPIO inputs and  text commands. You don’t have to go back to GPIO=0 to send text commands. You can send text  commands also in the GPIO=2, 3, 4, … modes. But when going from GPIO input to text commands, first  send a STOP command. Recommended maximum pulsing frequency: 20 kHz.

5.1. Pulse and direction mode

This mode is activated by the command GPIO=2 or 3.

GPIO=2 enables the input signals: ground (pin 2), pulse (13), direction (14), enable (15), index (16). GPIO=3 enables the input signals: ground (pin 2), pulse (13), forward (14), backward (15), index (16)

On each positive edge of the PULSE signal, the target position is incremented or decremented with a  particular step size. The default step size is 1 (1 encoder unit). This step size can be changed with the  command STPS, e.g. STPS=10.

The DIRECTION signal determines the direction: incrementing or decrementing the target position. The  ENABLE signal enables the input, it has to be high in order to start moving.

An alternative to the DIRECTION and ENABLE signal is the FORWARD and BACKWARD signal. By setting  one of them high, it’s possible to select the direction of moving.

The INDEX signal is an analog input used as digital input. When going high (positive edge) the controller  searches the index (3.3 V is sufficient, max. 10 V).

5.2. Encoder-like signal (A quad B input)

This mode is activated by the command GPIO=4

Input signals: ground (pin 2), A (13), B (14), enable (15), index (16).

Encoder-like A quad B signal used as input to adapt the target position.

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

5.3. PWM control

This mode is activated by the command GPIO=8 or 9

GPIO=8 enables the input signals: ground (pin 2), PWM (13), direction (14), enable (15). GPIO=9 enables the input signals: ground (pin 2), PWM (13), forward (14), backward (15).

The frequency of the PWM signal can be set by the command PWMF. PWMF=1000 sets the PWM  frequency to 1000 Hz. Default = 1000.

The speed is proportional to the pulse width. When the pulse width is 50 %, speed is set to 50 % of the  speed set by the SSPD command. When the signal is all the time high, then 100 % of SSPD is selected.  Speed can be controlled from 0 to 100 %.

The use of the direction, enable, forward and backward signals is the same as in pulse and direction  mode.

5.4. Analog controls

This mode is activated by the command GPIO=12 or 13

GPIO=12 enables the input signals: ground (pin 2), speed (1), direction (14), enable (15). GPIO=13 enables the input signals: ground (pin 2), speed (1), forward (14), backward (15).

The speed input is proportional to the voltage applied to the speed input pin. 10 V corresponds to 100  % of the speed set by the SSPD command.

The use of the direction, enable, forward and backward signals is the same as in pulse and direction  mode.

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

6. COMMUNICATION USING ETHERCAT

EtherCAT is an Ethernet-based high-performance industrial protocol used for multi-axis, high-speed  communication. Xeryon offers EtherCAT compatibility through a 4- or 16-axis motherboard, allowing  multiple actuators or stages to be controlled synchronously.

All manuals, technical drawings and 3D files and a TwinCAT 3 demo program can be found in our  EtherCAT package available for download from xeryon.com. Vendor ID: 0x0000004E

4-axis EtherCAT base plate

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

7. CONNECTIONS / PIN LAYOUTS

7.1. Connection to actuator

Actuator connector: ZIF 12 core

Pin 1 is defined by an arrow on the ZIF connector.

Connector: FH52E-12S-0.5SH

WARNING: Make sure to always use an opposing contact ZIF cable.

PIN

#SIGNAL FUNCTION IN/OUT PIN

#SIGNAL FUNCTION IN/OUT

1 ENC- Encoder Index - IN 7 ENC+ Encoder Index + IN 2 ENC B- Encoder B- IN 8 ENC GND Encoder ground IN 3 ENC A- Encoder A- IN 9 P1 Piezo phase 1 OUT 4 ENC PWR Encoder power OUT 10 P2 Piezo phase 2 OUT 5 ENC A+ Encoder A+ IN 11 P3 Piezo phase 3 OUT 6 ENC B+ Encoder B+ IN 12 P4 Piezo phase 4 OUT

7.2. Connection to stage

Stage connector: (14 pin, 0,1 inch pitch)

Connector: Molex 702471451

Pin 1 is defined by an arrow on the connector, see figure.

PIN

#SIGNAL FUNCTION IN/OUT PIN

#SIGNAL FUNCTION IN/OUT

1 ENC B- Encoder B- IN 8 SHIELD SHIELD IN 2 ENC B+ Encoder B+ IN 9 ENC PWR Encoder power OUT 3 ENC A- Encoder A- IN 10 ENC GND Encoder ground IN 4 ENC A+ Encoder A+ IN 11 P2 Piezo phase 2 OUT 5 ENC I- Encoder I- IN 12 P1 Piezo phase 1 OUT 6 ENC I+ Encoder I+ IN 13 P4 Piezo phase 4 OUT

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

7.3. Connection to GPIO

IO pins (0,05 inch pitch)

connector: CNC Tech 3220-26-0300-00

PIN

#SIGNAL FUNCTION IN/ OUT

PIN

#SIGNAL FUNCTION IN/ OUT

| AI0 | Analog speed input | IN | 14 | IN 2 | Direction / forward / B* |
| --- | --- | --- | --- | --- | --- |
| GND | Ground | IN | 15 | IN 3 | Enable / backward* |
| GND | Ground | IN | 16 | IN 4 | Find index |
| +5V | +5V | OUT | 17 | IN 5 | / |
| OUT PWR | Buffer supply voltage** | IN | 18 | IN 6 | UART Rx |
| +3,3V | +3,3V | OUT | 19 | / |  |
| OUT 1 | Trigger output | OUT | 20 | / |  |
| OUT 2 | Encoder valid flag | OUT | 21 | / |  |
| OUT 3 | Position reached flag | OUT | 22 | / |  |
| OUT 4 | Error flag | OUT | 23 | / |  |
| OUT 5 | Running flag | OUT | 24 | / |  |
| OUT 6 | UART Tx | OUT | 25 | GND | Ground |

1 IN

2 IN 3 IN 4

5IN

6

7

8

9

10

11

12 IN 13 IN 1 Pulse / A / PWM* IN 26 nSRST /

* For more details about these pins, see section “7 Communication using digital and analog IO” about the  GPIO. Digital signals are 3.3V logic and analog inputs are max 10V.

** NOTE: A voltage has to be supplied to this pin for UART to work. You can connect it to the available 3,3V  power pin or the 5V output pin. It is also possible to supply your own voltage (max 5V). The voltage of the  output pins will be determined by this pin (pin 5).

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

8. TUNING

Xeryon’s actuators are primarily intended for closed-loop control and for this reason already have a  position sensor integrated. A control algorithm is implemented on our controller and this algorithm  uses multiple control parameters. These control parameters have to be set according to the type of  actuators, load and customer-specific motion requirement. These parameters are pre-set by Xeryon

upon delivery of the actuators and controller. In most cases, you don’t need to adapt the settings right  out of the box. Nevertheless, it may be required that the user modifies these parameters for optimal  performance.

Below you can find the corresponding commands for the most common settings:

Below you can find the corresponding commands for the most common settings: Setting Command Setting Command

Shortest path

| PATH |  |
| --- | --- |
| ZON1 | Zone size, zone 2 |
| FREQ | Frequency, zone 2: |
| PROP | Proportional factor, zone 2 |
| PTOL | Speed |
| PTO2 | Accel / deceleration |
| TOUT | Mass |
| DUCO | Left soft limit |
| MIMP/MAMP | Right soft limit |

Zone size, zone 1 ZON2 Frequency, zone 1 FRQ2 Proportional factor, zone 1 PRO2 Position tolerance SSPD Position tolerance 2 ACCE/DECE

Timeout time MASS Amplitude control LLIM Min / max amplitude HLIM Phase correction PHAS Error limit ELIM

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

Frequencies/proportional values

Typically, in zone 1 (closest to the target position), the frequency and proportional factor are both  chosen higher. This gives a better ‘landing’ on the target position. For zone 2 (further away from the  target position) the frequency is chosen lower to increase speed. At the same time the proportional  factor for zone 2 typically has to be chosen lower to avoid instability. Be aware that outside a certain  frequency range, the motor will have very limited force (frequency too high) or feature unstable  behaviour (frequency too low). A typical frequency difference between FREQ and FRQ2 is between 1  and 3 kHz for a 1N motor and between 0.5 and 2 kHz for a 3N, 5N or 10N motor. The proportional  factors in zone 1 (PROP) are typically 2-3 times the value of the proportional factors in zone 2.

When the actuator does not want to land on the target position, despite optimising frequency and  proportional factor for zone 1, try to increase the positioning tolerances PTOL and PTO2. See the  instruction set for more information.

To obtain the maximum power out of the actuator, the frequency in zone 2 (FRQ2) should be set as low  as possible without drawing too much current, which results an amplifier fault (status bit 2 & 3).

Adding mass to the stage

As soon as adding a mass of 100g or more, the MASS parameter can be adapted. For every 100g of  mass you add, increase the MASS parameter with 100. Finding the optimal MASS parameters is usually  obtained by trial-and-error. This parameter is mostly used for heavy payloads on horizontal actuators.

Changing the dynamics

Would you like your actuator to react faster or slower, the following parameters can be adapted to  achieve this.

• SSPD is used to set the maximum speed the actuator. Please note, if the frequency in zone 2 is set  too high, the speed can’t always be reached.

• ACCE defines the acceleration of the actuator to the set speed in SSPD. The default value of 65500  means no acceleration limitation is taking into account (i.e. full acceleration).

• DECE defines the deceleration of the actuator upon reaching its target. The default value of 65500  means no deceleration limitation is taken into account (i.e. full deceleration).

• PROP & PRO2 are proportional factors used for the closed feedback loop. Both parameters are  used for the two different zones (ZON1 and ZON2). Typically PROP is higher than PRO2. Increase  them to get faster positioning times.

When the actuator needs to move with less overshoot, decrease the above mentioned parameters.  Higher proportional factors let the controller react stronger and reduce positioning errors, but can also  lead to instability or noisy operation when chosen too high.

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

9. WINDOWS GUI

To provide the user with a quick way to interact with the controlled and the connected actuator, a  Windows GUI is supplied with every controller. The use is simple and self-explanatory. It can be used  for manual input and to run simple scripts.

9.1. Required files

The Windows GUI makes use of the following files:

• The executable of the Windows GUI: Xeryon_Dialog.exe

• A configuration file named “config.txt”. This file should not be edited by the user. • A default settings file named “settings_default.txt”. The GUI reads this file for initial settings at start up. Replace or modify this file to alter the default settings. After saving, these values are stored on  the driver.

• A settings file named “settings_user.txt”. This file will be created by the program after pressing the  “save to file” button.

• A demo program. The file dialog window presents “demo.txt” as default filename. Remark: config.txt has to be in the same folder as Xeryon_Dialog.exe.

9.2. Commands for the Windows GUI demo program

Command Explanation

BAUD Set the baud rate for communication.

DPOL Delay used when polling for a ‘position reached’ signal after a new target position is set.  When DPOL is too small, the Windows GUI may trigger on the ‘position reached’ status  flag of the previous target position due to communication delay. In that case, a  succeeding WAIT command will start the timer at the start of the movement instead of  after the target has been reached.

HELP Switch help on or off. HELP=1 switches the info tips on. HELP=0 switches the info tips  off.

HALT Stop the program. (Not to be confused by the STOP command for the driver.) LABL Label in the program used by REPT.

LOG Start or stop logging of data. LOG=1 switches logging on. LOG=0 switched logging off.  Data is stored in datalog.csv. When datalog.csv already exists, new data is appended. MASS Specifies the mass/inertia of the load on the actuator. The Windows GUI calculates the  optimal control parameters to obtain stable operation.

MMAS Maximum mass that can be selected in the Windows GUI.

MPRO Maximum proportional factor that can be selected in the Windows GUI. MSPD Maximum speed that can be selected in the Windows GUI.

PORT Default port number to appear in the Windows GUI.

REPT Repeat the above program a specified number of times. The first argument specifies  the number of loops. The second argument specifies the label to jump to (label range 0  – 99). If the label does not exist, then the program jumps back to the first line. The REPT  command should be placed at the end of the block to be repeated. Nesting of REPT  blocks is allowed. Example: REPT=10 2 does 10 loops starting from label 2.

WAIT Wait a specified time before proceeding to the next command. Time expressed in  milliseconds. When WAIT follows a STEP or DPOS command, the timer is started when  reaching the target position.

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

Example program file (demo.txt)

SSPD=100 % Set speed to 100 mm/s or 100 degrees/s

DPOS=0 % Go to position 0

WAIT=100 % Wait 100 ms after arrival at position

LABL=2 % Set label 2

DPOS=60 % Go to position 60 mm or 60 degrees

WAIT=100 % Wait 100 ms after arrival at position

SSPD=10 % Set speed to 10 mm/s or 10 degrees/s

SCAN=1 % Move with constant speed in positive direction

WAIT=2000 % Wait for 2 s (while scan goes on)

REPT=3 2 % Repeat 3 times the code above, starting from label 2

STOP % Stop actuator

DPOS=0 % Finish in the centre

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

Function of buttons and status displays

1. Settings button:

− Load from file: select a file to load your own user settings.

− Edit: open the settings window (see below).

− Save to file: save the current settings to a new file.

− Save on driver: save the current to the controller.

2. Save to file: save the settings to a file you select

3. Connect to port: when the port is selected, connect to the port and select settings file 4. Port: Choose COM port of XD-OEM controller.

5. Connected to driver: successfully connected to the controller.

6. Communication error: communication error between the computer and controller. 7. Recover: recover the controller after a fault/error.

8. Reset encoder: reset encoder count. Index will need to be found again.

9. Find index: start indexing procedure.

10.Amplifier fault: overcurrent, undervoltage or short-circuit error.

11. Position error: position couldn’t be reached within set timeout.

12. Encoder error: encoder not working or damaged. (commonly caused by short-circuit) 13. Position reached: desired position reached.

14. Index found: index or home position found.

15. Stop: stop the motor.

16. Reset: reset the settings back to settings_default.txt and stop the motor 17. Start program: start the program selected below

18. Demo.txt: select the demo program

19. Blue slider: use the slider to change the desired position.

20. Scan + X: the actuator will continuously move in the + X direction.

21. Step + X: the actuator will take a step in the + X direction.

22. 0 button: move to the middle/index position.

23. Actual position: current position of the actuator. (can’t be changed here) 24. Desired position: change the position of the actuator here.

25. Step size: choose the size of the steps.

26. Actual speed: calculated speed of the actuator. (will always be lower than actual speed due to  polling delay)

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

10. PYTHON, MATLAB, C++ AND LABVIEW

Interaction with the XD-OEM controller and connected actuator is also possible through Python,  MATLAB, LabVIEW, C++ or any other programming language that supports sending ASCII commands  over a (virtual) COM port. Demo programs available for the programming languages mentioned above.  The demo’s include functions for advanced motion commands, error statuses, settings control and  more.

All libraries are available from our website under Downloads.

LabVIEW example

Last updated: 28/08/2025. All specifications are subject to change without prior notice.

11. SUPPORT

• Contact: support@xeryon.com

• Address: Interleuvenlaan 21, 3001 Leuven, Belgium • Phone: +32 (0)16 903 904

• Website: xeryon.com

Last updated: 28/08/2025. All specifications are subject to change without prior notice.